## Slide 1: Instrumentation Rapid Prototyping using LabVIEW and Arduino High-Precision Measurement Shield
Carmine Paolino (1156420)
Tutor: Prof. Chen Kangli
Tongji University - 15 June 2017

## Slide 2: The location
China Innovation Center
Focus on Biology, Chemistry and Engineering
R&D and training center
Created by Thermo Fisher Scientific
Prototypes are important
Need to build them fast and efficiently

## Slide 3: Some problems
Each measurement devices has its own interface
It takes time to configure a new platform
Time-to-prototype is critical for project development
Prototyping platforms may have limited performances

## Slide 4: Solution (1)
Common platform managing the devices
Intuitive graphical interface
Most widely used acquisition parameters
Implemented as a LabVIEW software
  Continuous Acquisition Framework

## Slide 5: Solution (2)
Extension boards can provide extended functionalities
Sigma delta ADCs allow to measure with up to 22 bits of resolution
PCB designed to improve Arduino microcontroller boards
  Arduino Measurement Shield

## Slide 6: Continuous acquisition framework
Forked from a National Instruments’ sample project
New devices and functionalities have been added
Some more to come

## Slide 7: Continuous acquisition framework > Hierarchy
Main.vi
Settings.vi
Acquisition Message Loop.vi
Keithley DMM 7510
PicoTech PicoScope 2000a series
and many more files…

## Slide 8: Main.vi > Block Diagram

## Slide 9: Main.vi > Abstract View
Event Handling loop
UI Message loop
Acquisition Message loop
Logging Message loop
Data Display loop
Event Handling Loop
UI Message
Loop
Acquisition Message Loop
Logging Message Loop
Data Display Loop
Data queue
Data notifier
UI queue
Acquisition
and UI queue
Logging
and UI queue

## Slide 10: Main.vi > Implementation Details
Notifier data structure
  Lossy
  Non-blocking
  Broadcast
Visual indicators managed by global variables
  Accessible anywhere
  Definition in front panel only

## Slide 11: Settings Dialogue.vi
Configured before acquiring
Common parameters
  Range
  Coupling
  Sampling Time
  ...
Device-dependent options
  Avoid configuration errors
  More work for the programmer

## Slide 12: Settings Dialogue.vi > DAQ Value Change Event
Changing DAQ Type triggers options update
Each device has its own
Need to specify them when adding a new one

## Slide 13: Settings Dialogue.vi > DAQ Value Change (2)

## Slide 14: Settings Dialogue.vi > OK Button Value Change

## Slide 15: Acquisition Message Loop.vi
Every state calls a sub-VI
Every sub-VI has a case structure
  switch (DAQ Type){…}
State = START
State = ACQUIRE
State = STOP

## Slide 16: DAQ Type 1 > Keithley DMM7510
Device specifications
SCPI interface
Configuration commands
Retrieval commands
Wrapping pseudocode
Extending the framework

## Slide 17: DMM7510 > Device specifications
7 ½ digits DMM
nV and pA sensitivity
1,000,000 readings/s
Touch screen display
USB connectivity
Support for SCPI commands

## Slide 18: DMM7510 > SCPI interface
Standards Commands for Programmable Instruments
Developed on top of the IEEE 488.2 standard
Tree-like structure
SENS1:VOLTage:DC:RANGe 10 // set DC voltage range to 10

## Slide 19: DMM7510 > Configuration commands
TRAC:MAKE “RPFbuff” 10000000 // create user buffer
TRAC:FILL:MODE CONT “RPFbuff” // wrap around when full
TRIG:TIM1:COUN 0; // restart Timer1 indefinitely
TRIG:TIM1:DEL <samplingTime>; // define overflow period
TRIG:TIM1:STIM COMM // start on software trigger
TRIG:TIM1:STAT ON // enable timer trigger state
SENS1:VOLT:DC:RANG <manualRange>; // setup DC voltage range
SENS1:VOLT:DC:APER <max(0.24, samplingTime)>; // select resolution
TRIG:MEAS:STIM TIM1 // sample on Timer1 overflow
VOLT:DC:AZERo:STATe OFF // disable the autozero feature
*TRG // send Software Trigger
Test.vi: configuration

## Slide 20: DMM7510 > Retrieval commands
TRAC:ACT:END? “RPFbuff” // get last element index
TRACE:DATA? <previousIndex+1>, <currentIndex>, “RPFbuff”, REL
/* Returns a string, which is a sequence as below
“<sample1_str>,<time1_str>,<sample2_str>,<time2_str>,…” */
Test.vi: acquisition

## Slide 21: DMM7510 > Wrapping pseudocode
What happens when wrapping around?
End index must be greater than start index
if Current_Index < Previous_Index then
Fetch_Buffer[Previous_Index + 1:Current_Index]
else
Fetch_Buffer[Previous_Index + 1:Buffer_Size]
Fetch_Buffer[1:Current_Index]
end

## Slide 22: DMM7510 > Extending the framework
Configure Hardware.vi > Configure DMM7510.vi
Acquire.vi > Generate DMM7510 Data.vi

## Slide 23: DAQ Type 2 > PicoTech PicoScope 2000a Series
Device specifications
LabVIEW streaming example
Accessing Memory buffers
Controlling the Timing
Calibrating the Device

## Slide 24: PS2000a > Device specifications
2 channels
100 MHz bandwidth
1 Gsps
8 kS to 128 MS memory capacity

## Slide 25: PS2000a > LabVIEW streaming example
C-language API
Reserved memory buffer
What happens when splitting configuration and acquisition in different subVIs?
PicoScope2000aExampleStreaming.vi

## Slide 26: PS2000a > Accessing Memory buffers
Values are copied into the cluster
The copy has a different address
Can’t get the samples anymore!
Configure Hardware.vi > PS2000a case

## Slide 27: PS2000a > Accessing Memory buffers (2)
---
Create reference
In-place structure to dereference
Equivalent to C pointers operations
What about the relative timing?
Updated example

## Slide 28: PS2000a > Controlling the Timing
Order of operation matters
Ad-hoc structure required to enforce it
Finally it works!
Definitive version

## Slide 29: PS2000a > Calibrating the Device
The device returns ADC counts (not Volts)
A sub-VI performs the conversion
It gives us access to the calibration constants
PS2000ADCtoVolts.vi

## Slide 30: 6 Arduino Measurement Shield PCB Design
Arduino
Design objectives
AD7176-2
Schematic – signal
Schematic – auxiliary
Placement & Routing
Layers
Updates
Serial interface
Communication reliability

## Slide 31: Arduino
Low cost microcontroller boards
8 bit @ 16 MHz up to 32 bit @ 84 MHz
5 V vs 3.3 V
14 vs 54 I/O lines
IDE and libraries
Shields
Arduino DUE
Arduino UNO

## Slide 32: Design objectives
16 noise-free bits
250 ksps
Signals up to 5 V
Can we make it?
AD7176-2: 250 ksps
AD7176-2: 5 sps

## Slide 33: AD7176-2
Sigma delta ADC
22 noise free bits at 5 sps
  or 16 bits at 250 ksps.
Single ended or differential configuration
Internal or external references and clock
SPI digital interface

## Slide 34: Schematic Diagram – Signal

## Slide 35: Schematic Diagram – Auxiliary

## Slide 36: Placement & routing
Complete placement
Submodules

## Slide 37: Layers
TOP
MidLayer1
BOTTOM
MidLayer2

## Slide 38: Updates > New revision
5 V reference dropout
  → use DC input
IOVDD noise
  → separate supplies through a new power stage
Noise introduced by jumpers
  → no gain selection on input buffers
Increased precision on temperature measurement
  → voltage divider on the sensor voltage to center the dynamic in [0; 1.1] V

## Slide 39: Serial interface
Only 8N1 or 8N2 mode on Arduino UNO
Samples are 16 bits
  → must be decomposed in 8 bit frames
Each sample becomes two packets
  → 2 x (start + payload + stop) = 20 bit in total
250 ksps sampling rate
  → 5 Mbaud required
BUT reliable communication only up to 1 Mbaud

## Slide 40: Communication reliability
1 Mbaud
2 Mbaud

## Slide 41: The culprit
ATMega16U2
Manages the serial to USB conversion
Has hardware USB peripheral
  But a lot of software is involved
    (https://github.com/arduino/)
Specialized firmware?
  Or may be just a buffer overrun?

## Slide 42: Conclusion
Both projects will evolve
Real case testing needed
Basic functionalities already guaranteed

## Slide 43: Questions?
…
…